Keysight

ME1000 RF Circuit Design

Lab 1

Low-Noise Amplifier Characterization Using Noise Figure Analyzer

This courseware product contains scholarly and technical information and is protected by copyright laws and international treaties. No part of this publication may be reproduced by any means, be it transmitted, transcribed, photocopied, stored in a retrieval system, or translated into any language in any form, without the prior written permission of Acehub Vista Sdn. Bhd.

The use of the courseware product and all other products developed and/or distributed by Acehub Vista Sdn. Bhd. are subject to the applicable License Agreement.

For further information, see the Courseware Product License Agreement.

##### Objective

- To measure the noise figure of a low-noise amplifier (LNA) via a noise figure analyzer (NFA)

##### Equipment Required

- Keysight N8973A Noise Figure Analyzer

- Keysight N4000A SNS Series Noise Source

- ME1000 RF Transceiver Kit (Receiver Unit)

##### Accessories Required

- TRM standard calibration kit

- 2 × SMA(m)-to-SMA(m) coaxial cable

- 1 × USB cable

- A PC running Microsoft® Windows® XP/Vista, pre-installed with the RF Trainer Control Panel software

IMPORTANT:

Turn off the training kit when not in use. The kit will turn off automatically when no mouse or keyboard action is detected for more than 10 minutes. Always ensure that the casing is grounded and the cover is latched up before powering up the device.

# Introduction

## Basic Equipment Setup

Figure 1 – General Equipment Configuration for LNA Measurement Using a NFA

# Noise Source and Instrument Calibration

Figure 2 – Calibrating the NFA

Record the LNA noise figure before performing calibration. Make the connection as shown in Figure 2 to calibrate the noise source, cables, PCB board, and instrument before measuring the LNA noise figure.

- Use the following settings:

Start frequency:		700 MHz

Stop frequency:  	900 MHz

Number of points:  	11

- Load the Excessive Noise Ratio (ENR) of the noise source into the instrument. This can be done via a diskette provided by the noise source manufacturer.

- Start the calibration process.

#### Exercises

- What is the noise figure before and after the calibration?

Noise figure before calibration = __________ dB

Noise figure after calibration = __________ dB

# Noise Figure Measurement

When the calibration is completed, make the connection as shown in Figure 1 to measure the noise figure of the LNA. Remove the Through PCB board.

- On your PC, launch the RF Trainer Control Panel. Select the RX Unit from the board selection and click Connect to RF Trainer. In the RX Unit Control area, click LNA Off/On to power up the LNA. You will see the LNA PCB LED light up on the trainer when it is powered on.

#### Exercises

- Plot the gain and noise figure of the LNA from 700 MHz to 900 MHz.

- What is the noise figure and gain at 868 MHz?

Noise figure at 868 MHz = __________ dB

Gain at 868 MHz = __________ dB

###### References

- Application Note 1354, “Practical Noise-Figure Measurement and Analysis for Low-Noise Amplifier Designs,” Keysight Technologies

- Thomas H. Lee, “Planar Microwave Engineering,” Cambridge University Press, 2004

- David M. Pozar, “Microwave Engineering,” 3rd Edition, John Wiley & Sons, 2005