ME1000 RF Circuit Design

Lab 6

Frequency Synthesizer Characterization Using Spectrum Analyzer

This courseware product contains scholarly and technical information and is protected by copyright laws and international treaties. No part of this publication may be reproduced by any means, be it transmitted, transcribed, photocopied, stored in a retrieval system, or translated into any language in any form, without the prior written permission of Acehub Vista Sdn. Bhd.

The use of the courseware product and all other products developed and/or distributed by Acehub Vista Sdn. Bhd. are subject to the applicable License Agreement.

For further information, see the Courseware Product License Agreement.

##### Objective

- To measure the output power level, harmonic frequencies, and phase noise of a frequency synthesizer.

##### Equipment Required

- Keysight N9320B 3 GHz RF Spectrum Analyzer

- ME1000 RF Transceiver Kit (Receiver Unit)

##### Accessories Required

- 2 × SMA(m)-to-SMA(m) coaxial cable

- 1 × USB cable

- A PC running Microsoft® Windows XP/Vista®, pre-installed with the RF Trainer Control Panel software

IMPORTANT:

Turn off the training kit when not in use. The kit will turn off automatically when no mouse or keyboard action is detected for more than 10 minutes. Always ensure that the casing is grounded and the cover is latched up before powering up the device.

# Introduction

## Basic Equipment Setup

Figure 1 – General Equipment Configuration for Frequency Synthesizer Measurement Using a Spectrum Analyzer

# Determine the Local Oscillator (LO) Input Level Using a Spectrum Analyzer

- Connect the frequency synthesizer directly to the spectrum analyzer (SA), as shown in Figure 1.

- On your PC, launch the RF Trainer Control Panel. Select the RX Unit from the board selection and click Connect to RF Trainer. In the Frequency Synthesizer Control area, click Frequency Synthesizer Off/On to power up the synthesizer. You will see the frequency synthesizer PCB LED light up on the trainer when it is powered on. Set the frequency to 818 MHz.

- Set the following settings for the SA:

Centre frequency:  	818 MHz

Span:  			10 MHz

RBW:  			100 kHz

Attenuation:  		40 dB

N9320B setting: 		“[ ]”: Hardkey; “{ }”: Softkey

Preset to default settings: 	[Preset/System] > {Preset}

Center frequency: 		[FREQUENCY] > [818] > {MHZ}

Span: 				[SPAN] > [10] > {MHZ}

Attenuation: 			[AMPLITUDE] > {Attenuation} > [40] > {dB}

RBW: 				[BW/AVG] > {Res BW} > {100} > {KHZ}

#### Exercises

- What is the power level displayed on the SA?

LO level measured at SA, Psa_LO @818 MHz = __________ dBm

Note: Use the marker function Search Peak if necessary.

- Determine the actual output power level of the LO.

Actual LO output power, Plo 	= Psa_LO + Lcable

= __________ dBm + __________ dB

= __________ dBm

- Change the Span setting to Full Span on the SA to look at the harmonic tones.

N9320B setting: 	“[ ]”: Hardkey; “{ }”: Softkey

Span: 			[SPAN] > {Full Span}

Attenuation: 		[AMPLITUDE] > {Attenuation} > [10] or {auto} > {dB}

#### Exercise

- Sketch the harmonic tones and measure their respective power level.

# Phase Noise Measurement

- Use the same connection as shown in Figure 1.

- Use the following settings for the SA:

Centre frequency:  	818 MHz

Span:  			300 kHz

Input attenuation:  	20 dB (or auto)

RBW:  			10 kHz (or auto)

N9320B setting: 	“[ ]”: Hardkey; “{ }”: Softkey

Center frequency: 	[FREQUENCY] > [818] > {MHZ}

Span: 			[SPAN] > [300] > {KHZ}

Attenuation: 		[AMPLITUDE] > {Attenuation} > [20] > {dB}

RBW: 			[BW/AVG] > {Res BW} > {10} > {KHZ}

#### Exercise

- What is the power level displayed on the SA?

LO level measured at SA, Psa_LO @818 MHz = __________ dBm

Note:
Use the marker function Search Peak if necessary.

- Use the SA built-in noise marker to measure the phase noise:

Marker function: 	Phase noise on

Offset:  			Manual

Offset value:  		100 kHz, 50 kHz, and 10 kHz

N9320B setting: 			“[ ]”: Hardkey; “{ }”: Softkey

Place Marker @818 MHz: 		[Peak Search]

Marker: 	[Marker] > {Function} > {Phase Noise} > {Offset Manual} > [100] > {KHz}

Turn on Phase Noise Measurement: 	{Phase Noise On}

#### Exercises

- What is the phase noise (with 100 kHz offset) displayed on the SA?

Phase noise @100 kHz offset = __________ dBc/Hz

- What is the phase noise (with 50 kHz offset) displayed on the SA?

Phase noise @50 kHz offset = __________ dBc/Hz

- What is the phase noise (with 10 kHz offset) displayed on the SA?

Phase noise @10 kHz offset = __________ dBc/Hz

###### Reference

- Presentation slides, “A Seminar on RF Measurement – Spectrum Analysis Basics,” Keysight Technologies, 2001