ME1000 RF Circuit Design

Lab 2

Power Amplifier Characterization

Using RF Analyzer

This courseware product contains scholarly and technical information and is protected by copyright laws and international treaties. No part of this publication may be reproduced by any means, be it transmitted, transcribed, photocopied, stored in a retrieval system, or translated into any language in any form, without the prior written permission of Acehub Vista Sdn. Bhd.

The use of the courseware product and all other products developed and/or distributed by Acehub Vista Sdn. Bhd. are subject to the applicable License Agreement.

For further information, see the Courseware Product License Agreement.

Objectives

- To characterize a power amplifier (PA) with reflection and transmission measurement using an RF analyzer (RFA)

- To observe gain compression in the PA using the RFA

Equipment Required

- ME1000 Transmitter unit

- Keysight N9912A Fieldfox RF Analyzer (with the NA option)

Accessories Required

- 2 x SMA(m)-to-SMA(m) coaxial cable

- 1 x SMA(m)-to-SMA(m) jumper cable, 0.18 m

- 1 x RF power combiner

- 1 x TRM standard calibration kit

- 1 x USB cable

- A PC running Microsoft® Windows XP/Vista® with a minimum of 512 MB RAM, and pre-installed with the RF Trainer Control Panel software (Refer to the ME1000 Quick Start Guide for the installation procedure)

IMPORTANT:

Turn off the training kit when not in use. The kit will turn off automatically when no mouse or keyboard action is detected for more than 10 minutes. Always ensure that the casing is grounded and the cover is latched up before powering up the device.

# Introduction

Basic Equipment Setup

RF Trainer Control Panel

PC

USB cable

RF bandpass filter

Power

amplifier

Upconverter

Frequency synthesizer

RF Out

LO In

RF In

IF Out

In

Out

Out

In

USB port

Transmitter unit

SMA cable (keep this as short as possible)

SMA cable

RF Analyzer

RF OUT

RF IN

Figure 1 – General Equipment Configuration for PA Measurement Using the RFA

In measuring a high gain amplifier, it is possible to overload the amplifier input. This will cause the unit to saturate and produce a significant non-linearity in the output signal which will possibly damage the unit. It can also damage the test equipment due to high power amplification. You should either add more attenuation to the input or output of the amplifier or reduce the RF power of the signal source before repeating the measurement.

The RF OUT port of the RFA generates a sinusoidal wave. The wave interacts with the device-under-test (DUT) and the result is captured by the RF IN port of the RFA. Thus RF OUT is always connected to the input of the DUT and RF IN is connected to the output of the DUT.

In this lab, you need to set the RFA to the Network Analyzer (NA) mode. When performing a transmission or an S21 measurement, carry out the calibration procedures from Lab 1 with the SMA cables on both the RF OUT and RF IN ports of the RFA before you proceed. The cable and connectors on the RF OUT port of the RFA should be as short as possible and of high quality.

Note:

If your RFA comes with Option 111 (‘QuickCal’), you will be able to perform external calibration on the RF OUT port for reflection or an S11 measurement. Refer to the Keysight N9912A User’s Guide for the detailed procedure.

When viewing signals, the [Freq/Dist], [Scale/Amptd], and [Marker] hardkeys are the most helpful. When the [Scale/Amptd] hardkey is pressed, the {Autoscale} option can be quite handy, but you will rarely use it as you get more comfortable with manipulating the display.

Note:

Use the marker ([Marker] hardkey) to record your readings. The marker is used to obtain a reading at a particular frequency or amplitude. For simplicity, you may wish to use the Marker Function or Marker Search option to search for maximum and minimum points.

# Transmission Measurements

- Set the RFA (N9912A) to the NA mode for S21 measurement.

- Perform the ‘Normalization’ calibration as in Lab 1 from 2 MHz to 4 GHz with the LOW output power setting.

- Connect the PA to the N9912A as shown in Figure 1.

- Launch the RF Trainer Control Panel software from the PC. Select TX Unit from the board selection and click Connect to RF Trainer. In the TX Unit Control area, click Power Amplifier Off to power up the PA. You will observe the PA PCB LED light up on the training kit, and the button text change to Power Amplifier On.

- Use the following settings to determine the small-signal gain of the PA for a frequency range of within 700 MHz and 1 GHz, with the output power level of the N9912A set to LOW:

Start		:  700 MHz

Stop		:  1 GHz

Format 	:  Log magnitude

Measurement	:  S21

Power Level	:  LOW (corresponds to –20 dBm)

Scale		:  Autoscale

N9912A setting: “[ ]”: Hardkey; “{ }”: Softkey

Set to the NA mode for S21 measurement: [Mode] > {NA} > {S21}

Set the Start Frequency to 700 MHz: [Freq/Dist] > {Start} > [700] > {MHz}

Set the Stop Frequency to 1 GHz: [Freq/Dist] > {Stop} > [1] > {GHz}

Set the output power level to LOW: [Meas Setup 4] > {Output Power} > {Low}

Scale: [Scale/Amptd] > {Autoscale}

Perform Calibration: Normalization

Exercises

Observations/data recording

- Find the maximum gain and the frequency at which it occurs.

Gmax = _____ dB at ______ MHz

- Find the frequency below the maximum gain at which the gain is reduced by 3 dB from the maximum gain.

Gmax –3 dB occurs at _______ MHz

- Find the gain at 868 MHz.

Gain at 868 MHz = _________ dB

- Reverse the connections on the LNA to determine the isolation, i.e. connect the RF OUT port to the OUT terminal of the LNA and the RF IN port to the IN terminal of the LNA. Use Autoscale if necessary to view the complete waveform. This waveform is the S12 waveform since we are now sending the sinusoidal wave into the LNA output and sample its response at the LNA input instead.

Exercises

Observations/data recording

- Find the isolation at 868 MHz.

Isolation at 868 MHz = _________ dB

[Note: Use averaging, smoothing, and smoothing aperture if necessary]

- Elaborate the importance of isolation in an amplifier.

# Reflection Measurements

- Reconnect the PA to the N9912A as shown in Figure 1.

- Set the N9912A to the NA mode for S11 measurement. Make sure that the output power level is set to LOW. Keep the SMA cable at the RF OUT port of the N9912A as short as possible.

- Use the following settings to determine the insertion loss of the filter:

Format 	:  Log magnitude

Measurement	:  S11

Scale		:  Autoscale

N9912A setting: “[ ]”: Hardkey; “{ }”: Softkey

Set to the NA mode for S21 measurement: [Mode] > {NA} > {S11}

Set the display to the Log magnitude format: [Meas Setup] > {Format} > {LogMag}

Scale: [Scale/Amptd] > {Autoscale }

Exercises

Observations/data recording

- Find the minimum input return loss and the frequency at which it occurs.

Min return loss = _____ dB at ______ MHz

- Find the maximum input return loss and the frequency at which it occurs.

Max return loss = _____ dB at ______ MHz

- Find the input return loss at 868 MHz.

Return loss at 868 MHz = _________ dB

[Note: Use the marker function to search for maximum and minimum points]

- Reverse the connections on the PA to determine the output return loss, i.e. connect the RF OUT port of the N9912A to the OUT terminal of the PA and the RF IN port of the N9912A to the IN terminal of the PA. This waveform is the S22 waveform since we are now sending the sinusoidal wave into the PA output and sample its reflection with the terminated IN terminal.

- Repeat steps 3 to 5 to characterize the S22 in terms of its output return loss.

Exercises

Observations/data recording

- Find the minimum output return loss and the frequency at which it occurs.

Min return loss = _____ dB at ______ MHz

- Find the maximum output return loss and the frequency at which it occurs.

Max return loss = _____ dB at ______ MHz

- Find the output return loss at 868 MHz.

Return loss at 868 MHz = _________ dB

[Note: Use the marker function to search for maximum and minimum points]

# Gain Compression Observation

- Record the power gain of the PA at 868 MHz at low input power by following the procedure in Section 2, Transmission Measurements. When the N9912A output power level is set to LOW, the nominal output power is approximately –20 dBm into a 50  load.

- Next, set the N9912A output power level to HIGH:

Start		:  700 MHz

Stop		:  1 GHz

Format 	:  Log magnitude

Measurement	:  S21

Power Level	:  HIGH (corresponds to +3 dBm)

Scale		:  Autoscale

N9912A setting: “[ ]”: Hardkey; “{ }”: Softkey

Set to the NA mode for S21 measurement: [Mode] > {NA} > {S21}

Set the Start Frequency to 700 MHz: [Freq/Dist] > {Start} > [700] > {MHz}

Set the Stop Frequency to 1 GHz: [Freq/Dist] > {Stop} > [1] > {GHz}

Set the output power level to HIGH: [Meas Setup 4] > {Output Power} > {High}

Scale: [Scale/Amptd] > {Autoscale}

Perform Calibration: Normalization

- Record the power gain of the PA at 868 MHz at high input power; this corresponds to the N9912A RF OUT port providing a nominal power of +3 dBm to the PA.

- Finally, we would also like to record the power gain of the PA at 868 MHz at the medium input power level. To do this, we connect a 6 dB attenuator between the N9912A RF OUT port and the PA In port as shown in Figure 2. We can improvise a 6 dB attenuator using the power splitter PCB and the matched load PCB provided with this training kit.

PC

USB cable

RF bandpass filter

Power

amplifier

Upconverter

Frequency synthesizer

RF Out

LO In

RF In

IF Out

In

Out

Out

In

USB port

Transmitter unit

SMA cable (keep this as short as possible)

SMA cable

Power

splitter

Matched load

6 dB attenuator

RFA (N9912A)

RF OUT

RF IN

Figure 2 – Adding an Attenuator to Reduce the Input Power to the PA for Gain Compression Measurement Using the RFA (N9912A)

Exercises

Observations/data recording

- Discuss the behavior of the PA with respect to different input power levels in terms of  power gain.

Power gain at small input power level = _______ dB

Power gain at high input power level = _______ dB

- Define and elaborate the meaning of gain compression point. What is the implication of gain compression to the harmonic levels at the PA output?

References

- Application Note 1287-1, “Understanding the Fundamental Principles of Vector Network Analysis”, Keysight Technologies.

- Application Note 1287-2, “Exploring the Architectures of Network Analyzers”, Keysight Technologies.

- Application Note 1287-3, “Applying Error Correction to Network Analyzer Measurements”, Keysight Technologies.

- Application Note 1287-4, “Network Analyzer Measurements: Filter and Amplifier Examples”, Keysight Technologies.

- Product Note 8753-1, “RF Component Measurements: Amplifier Measurements Using the Keysight 8753 Network Analyzer”, Keysight Technologies.

- Thomas H. Lee, “Planar Microwave Engineering”, Cambridge University Press, 2004.