ME1000 RF Circuit Design

Lab 4

Filter Characterization Using RF Analyzer

This courseware product contains scholarly and technical information and is protected by copyright laws and international treaties. No part of this publication may be reproduced by any means, be it transmitted, transcribed, photocopied, stored in a retrieval system, or translated into any language in any form, without the prior written permission of Acehub Vista Sdn. Bhd.

The use of the courseware product and all other products developed and/or distributed by Acehub Vista Sdn. Bhd. are subject to the applicable License Agreement.

For further information, see the Courseware Product License Agreement.

Objectives

- To characterize the RF filter with reflection and transmission measurement via an RF analyzer (RFA)

- To display results in multiple plots such as the Smith Chart and Magnitude Plot

Equipment Required

- ME1000 Transmitter unit

- Keysight N9912A FieldFox RF Analyzer

Accessories Required

- 2 x SMA(m)-to-SMA(m) coaxial cable

IMPORTANT:

Turn off the training kit when not in use. The kit will turn off automatically when no mouse or keyboard action is detected for more than 10 minutes. Always ensure that the casing is grounded and the cover is latched up before powering up the device.

# Introduction

Basic Equipment Setup

RF Analyzer

RF OUT

RF IN

RF bandpass filter

Power

amplifier

Upconverter

Frequency synthesizer

RF Out

LO In

RF In

IF Out

In

Out

Out

In

USB port

Transmitter unit

SMA cable

SMA cable

Figure 1 – General Equipment Configuration for Filter Measurement

The RF OUT port of the RF analyzer (RFA) generates a sinusoidal wave. The wave interacts with the device-under-test (DUT) and the result is captured by the RF IN port of the RFA. Therefore, the RF OUT port is always connected to the input of the DUT and the RF IN port is connected to the output of the DUT.

In this lab, you need to set the RFA to the Network Analyzer (NA) mode. When performing a transmission, or an S21 measurement, carry out the calibration procedures from Lab 1 with the SMA coaxial cables connected to both the RF OUT port and RF IN port of the RFA before you proceed. The cable and connectors on the RF OUT port of the RFA should be as short as possible and of high quality.

Note:

If your RFA comes with Option 111 (‘QuickCal’), you will be able to perform external calibration on the RF OUT port for reflection or an S11 measurement. Refer to the Keysight N9912A User’s Guide for the detailed procedure.

When viewing signals, the [Freq/Dist], [Scale/Amptd], and [Marker] hardkeys are the most helpful. When the [Scale/Amptd] hardkey is pressed, the {Autoscale} option can be quite handy, but you will rarely use it as you get more comfortable with manipulating the display.

Note: 
Use the marker ([Marker] hardkey) to record your readings. The marker is used to obtain a reading at a particular frequency or amplitude. For simplicity, you may wish to use the Marker Function or Marker Search option to search for maximum and minimum points.

# Transmission Measurements

- Set the RFA (N9912A) to the NA mode for S21 measurement.

- Perform the Normalize calibration as in Lab 1 from 2 MHz to 4 GHz with the LOW output power setting.

- Use the following settings to determine the insertion loss of the filter:

Format 	:  Log magnitude

Measurement	:  S21

Scale		:  Autoscale

N9912A setting: “[ ]”: Hardkey; “{ }”: Softkey

Set to the NA mode for S21 measurement: [Mode] > {NA} > {S21}

Set the Start Frequency to 2 MHz: [Freq/Dist] > {Start} > [2] > {MHz}

Set the Stop Frequency to 4 GHz: [Freq/Dist] > {Stop} > [4] > {GHz}

Set the output power level to LOW: [Meas Setup 4] > {Output Power} > {Low}

Scale: [Scale/Amptd] > {Autoscale}

Perform Calibration: Normalization

- Connect the RF bandpass filter as shown in Figure 1.

Exercises

Observations/data recording – General

- What is the filter type? (bandpass, low pass, high pass, or not sure)

- Use the Marker function to determine the following parameters:

3-dB cut-off frequencies (use minimum loss as the reference value):

Lower 3-dB frequency (fL) = __________ MHz

Upper 3-dB frequency (fH) = __________ MHz

3-dB bandwidth = __________ MHz

Center frequency (fc) = ____________ MHz

[Note: The center frequency is defined as the mean of the lower and upper 3-dB frequencies.]

Insertion loss (S21) at the center frequency = ___________ dB

- What is the importance of insertion loss?

- What is the importance of insertion loss flatness?

- Explain why the S21 magnitude increases slowly in the upper stopband with frequency.

Observations/data recording – Out-of-band rejection

- Find the rejection at stopband (at 300 MHz, 500 MHz, and 800 MHz offset from the center frequency).

Upper rejection at 300 MHz = _____	dB

Lower rejection at 300 MHz = _____	dB

Upper rejection at 500 MHz = _____	dB

Lower rejection at 500 MHz = _____	dB

Upper rejection at 800 MHz = _____	dB

Lower rejection at 800 MHz = _____	dB

# Reflection Measurements

- Set the N9912A to the NA mode for S11 measurement.

- Use the following settings to determine the insertion loss of the filter:

Format 	:  Log magnitude

Measurement	:  S11

Scale		:  Autoscale

N9912A setting: “[ ]”: Hardkey; “{ }”: Softkey

Set to the NA mode for S21 measurement: [Mode] > {NA} > {S11}

Set the display to the Log magnitude format: [Meas Setup] > {Format} > {LogMag}

Scale: [Scale/Amptd] > {Auto Scale}

Exercises

Observations/data recording

  - What is the return loss at passband?

The return loss, S11 at passband = ________dB

  - What is the return loss at stopband?

The return loss, S11 at stopband = ________dB

[Note: You can choose any frequency points within passband or stopband]

  - What is the importance of the return loss at passband?

  - From the return loss observed at passband and stopband, describe the filter matching condition at both bands.

- Change the settings to display the Smith Chart for impedance measurement.

- From the Smith Chart, you can determine the magnitude and phase of the reflection coefficient. From these values, you can convert the reflection coefficient into the complex form, and finally convert the reflection coefficient into the input impedance Zin, 
via Zin=1+Γ111−Γ11⋅Zo , where Zo is 50 

Format 	:  Smith Chart

N9912A setting: “[ ]”: Hardkey; “{ }”: Softkey

Set the display to the Smith Chart format: [Meas Setup] > {Format} > {Smith}

Exercises

Observations/data recording

- What is the impedance at passband?

The impedance at passband = ( ______ + j _______ ) Ω

- What is the impedance at stopband?

The impedance at stopband = ( ______ + j _______ ) Ω

[Note: You can choose any frequency points within passband or stopband]

- What is the ideal impedance at passband?

The ideal impedance at passband = ( ______+ j _______ ) Ω

- What is the ideal impedance at stopband?

The ideal impedance at stopband = ( ______+ j _______ ) Ω

References

- Application Note 1287-1, “Understanding the Fundamental Principles of Vector Network Analysis”, Keysight Technologies.

- Application Note 1287-2, “Exploring the Architectures of Network Analyzers”, Keysight Technologies.

- Application Note 1287-4, “Network Analyzer Measurements: Filter and Amplifier Examples”, Keysight Technologies.

- “Keysight FieldFox RF Analyzer N9912A User’s Guide”, Keysight Technologies, 2009.

- Thomas H. Lee, “Planar Microwave Engineering”, Cambridge University Press, 2004.