ME1000 RF Circuit Design

Lab 5

Mixer Characterization Using

RF Analyzer

This courseware product contains scholarly and technical information and is protected by copyright laws and international treaties. No part of this publication may be reproduced by any means, be it transmitted, transcribed, photocopied, stored in a retrieval system, or translated into any language in any form, without the prior written permission of Acehub Vista Sdn. Bhd.

The use of the courseware product and all other products developed and/or distributed by Acehub Vista Sdn. Bhd. are subject to the applicable License Agreement.

For further information, see the Courseware Product License Agreement.

Objective

- To measure SWR, return loss, RF feedthrough, and LO isolation of a mixer using an RF analyzer (RFA)

Equipment Required

- ME1000 Receiver unit

- Keysight N9912A Fieldfox RF Analyzer

Accessories Required

- 1 x TRM standard calibration kit

- 2 x SMA(m)-to-SMA(m) coaxial cable,

- 1 x SMA(m)-to-SMA(m) jumper cable,

- 1 x USB cable

- A PC running Microsoft® Windows XP/Vista® with a minimum of 512 MB RAM, and pre-installed with the RF Trainer Control Panel software (Refer to the ME1000 Quick Start Guide for the installation procedure)

IMPORTANT:

Turn off the training kit when not in use. The kit will turn off automatically when no mouse or keyboard action is detected for more than 10 minutes. Always ensure that the casing is grounded and the cover is latched up before powering up the device.

# Introduction

Basic Equipment Setup

RF OUT

SMA cable

SMA cable

PC

RF Trainer Control Panel

USB cable

USB port

SMA jumper cable

Receiver unit

To Mixer Output (IF)

To Mixer Input (RF) (keep the cable as short as

possible)

Synthesizer output

RF Analyzer

RF IN

LNA

Down

mixer

Frequency synthesizer

IF bandpass

filter

IF amplifier

LO In

RF Out

RF In

IF Out

Out

In

In

In

Out

Out

Figure 1 – General Equipment Configuration for Mixer Measurement Using the RFA

The RF OUT port of the RFA generates a sinusoidal wave. The wave interacts with the device-under-test (DUT) and the result is captured by the RF IN port of the RFA. Thus RF OUT is always connected to the input of the DUT and RF IN is connected to the output of the DUT.

In this lab, you need to set the RFA to the Network Analyzer (NA) mode. When performing a transmission or an S21 measurement, carry out the calibration procedures from Lab 1 with the SMA cables on both the RF OUT and RF IN ports of the RFA before you proceed. The cable and connectors on the RF OUT port of the RFA should be as short as possible and of high quality.

Note:

If your RFA comes with Option 111 (‘QuickCal’), you will be able to perform external calibration on the RF OUT port for reflection or an S11 measurement. Refer to the Keysight N9912A User’s Guide for the detailed procedure.

When viewing signals, the [Freq/Dist], [Scale/Amptd], and [Marker] hardkeys are the most helpful. When the [Scale/Amptd] hardkey is pressed, the {Autoscale} option can be quite handy, but you will rarely use it as you get more comfortable with manipulating the display.

Note:

Use the marker ([Marker] hardkey) to record your readings. The marker is used to obtain a reading at a particular frequency or amplitude. For simplicity, you may wish to use the Marker Function or Marker Search option to search for maximum and minimum points.

# RF/IF Port SWR Measurement with the RFA

PC

RF Trainer Control Panel

USB cable

SMA jumper cable

Synthesizer output

SMA adapter

SMA cable

Mixer input (RF)

Matched load

RFA (N9912A)

RF OUT

RF IN

USB port

LNA

Down

mixer

Frequency synthesizer

IF bandpass

filter

IF amplifier

LO In

RF Out

RF In

IF Out

Out

In

In

In

Out

Out

Figure 2 – RF Port SWR Measurement with the RFA

Make the connection as shown in  to determine the RF In port SWR. Note that you must connect the synthesizer output into the LO In port of the mixer via a coaxial jumper cable. Also, terminate the IF Out port of the mixer with the 50 Ω matched PCB board, since the mixer is designed to drive a 50  load at its IF Out port. The RF IN port of the N9912A is not used. Keep the cable at the RF OUT port of the  N9912A as short as possible.

- Set the N9912A to the NA mode for S11 measurement. Make sure that the output power level is set to LOW.

- Launch the RF Trainer Control Panel software from the PC. Select RX Unit from the board selection and click Connect to RF Trainer. In the RX Unit Control area, click Mixer Off to power up the mixer. You will observe the mixer PCB LED light up on the training kit, and the button text change to Mixer On. Next, in the Frequency Synthesizer Control area, click Frequency Synthesizer Off to power up the synthesizer and set the output frequency to 818.0 MHz.

- Use the following settings to determine the SWR of the RF In port of the mixer:

Start		:  700 MHz

Stop		:  900 MHz

Format 	:  VSWR (Voltage standing wave ratio)

Measurement	:  S11

Scale		:  Autoscale

N9912A setting: “[ ]”: Hardkey; “{ }”: Softkey

Set to the NA mode for S11 measurement: [Mode] > {NA} > {S11}

Set the Start Frequency to 700 MHz: [Freq/Dist] > {Start} > [700] > {MHz}

Set the Stop Frequency to 900 MHz: [Freq/Dist] > {Stop} > [900] > {MHz}

Set the output power level to LOW: [Meas Setup 4] > {Output Power} > {Low}

Set the display to the VSWR format: [Meas Setup] > {Format} > {VSWR}

Scale: [Scale/Amptd] > {Autoscale}

Exercise

Observations/data recording

- What is the voltage standing wave ratio at 868 MHz?

SWR at 868 MHz = 1: _______

- Change the setting to display the log magnitude.

Format 	:  Log magnitude

N9912A setting: “[ ]”: Hardkey; “{ }”: Softkey

Set the display to the log magnitude format: [Meas Setup] > {Format} > {Log Mag}

Scale: [Scale/Amptd] > {Autoscale}

Exercises

Observations/data recording

- What is the input return loss at 868 MHz?

Input return loss at 868 MHz = _______dB

- Plot the input return loss from 700 MHz to 900 MHz.

Freq (MHz)

Return loss, S11 (dB)

- Connect the IF Out port to the RF IN port of the N9912A and terminate the mixer RF port with the matched load. Change the start and stop frequencies to 10 MHz and 100 MHz respectively. Make sure that you turn off the mixer DC supply before removing any connections by clicking Mixer Off in the RF Trainer Control Panel software. Power on the mixer again when the proper connection has been made.

Start		:  10 MHz

Stop		:  100 MHz

Format 	:  VSWR

Scale		:  Autoscale

N9912A setting: “[ ]”: Hardkey; “{ }”: Softkey

Set the Start Frequency to 10 MHz: [Freq/Dist] > {Start} > [10] > {MHz}

Set the Stop Frequency to 100 MHz: [Freq/Dist] > {Stop} > [100] > {MHz}

Set the output power level to LOW: [Meas Setup 4] > {Output Power} > {Low}

Set the display to the VSWR format: [Meas Setup] > {Format} > {VSWR}

Scale: [Scale/Amptd] > {Autoscale}

Exercise

Observations/data recording

- What is the standing wave ratio at 50 MHz?

SWR at 50 MHz = 1: _______

- Change the setting to display the log magnitude.

Format 	:  Log magnitude

N9912A setting: “[ ]”: Hardkey; “{ }”: Softkey

Set the display to the log magnitude format: [Meas Setup] > {Format} > {Log Mag}

Scale: [Scale/Amptd] > {Autoscale}

Exercises

Observations/data recording

- What is the input return loss at 50 MHz?

Input return loss at 50 MHz = _______ dB

- Plot the input return loss from 10 MHz to 100 MHz.

Freq (MHz)

Return loss, S11 (dB)

# LO Port SWR Measurement with the RFA

PC

RF Trainer Control Panel

USB cable

USB port

Receiver unit

SMA adapter

Matched load

SMA adapter

SMA cable

LNA

Down

mixer

Frequency synthesizer

IF bandpass

filter

IF amplifier

LO In

RF Out

RF In

IF Out

Out

In

In

In

Out

Out

RFA (N9912A)

RF OUT

RF IN

Figure 3 – LO Port SWR Measurement with the RFA

Make the connection as shown in Figure 3 to determine the LO port SWR. Note that you need to terminate both the IF Out and RF In ports of the mixer with a 50 Ω matched PCB board. The frequency synthesizer is not used in this measurement. Keep the cable at the RF OUT port of the N9912A as short as possible.

- Launch the RF Trainer Control Panel software from the PC. Select RX Unit from the board selection and click Connect to RF Trainer. In the RX Unit Control area, click Mixer Off to power up the mixer. You will observe the mixer PCB LED light up on the training kit, and the button text change to Mixer On.

- Use the following settings to determine the SWR of the RF port:

Start		:  700 MHz

Stop		:  900 MHz

Format 	:  SWR

Measurement	:  S11

Scale		:  Autoscale

N9912A setting: “[ ]”: Hardkey; “{ }”: Softkey

Set to the NA mode for S11 measurement: [Mode] > {NA} > {S11}

Set the Start Frequency to 700 MHz: [Freq/Dist] > {Start} > [700] > {MHz}

Set the Stop Frequency to 900 MHz: [Freq/Dist] > {Stop} > [900] > {MHz}

Set the output power level to LOW: [Meas Setup 4] > {Output Power} > {Low}

Set the display to the VSWR format: [Meas Setup] > {Format} > {VSWR}

Scale: [Scale/Amptd] > {Autoscale}

Exercise

Observations/data recording

- What is the standing wave ratio at 818 MHz?

SWR at 818 MHz = 1: _______

- Change the setting to display the log magnitude plot.

Format 	:  Log magnitude

N9912A setting: “[ ]”: Hardkey; “{ }”: Softkey

Set the display to the log magnitude format: [Meas Setup] > {Format} > {Log Mag}

Scale: [Scale/Amptd] > {Autoscale}

Exercises

Observations/data recording

- What is the input return loss at 818 MHz?

Input return loss at 818 MHz = _______ dB

- Plot the input return loss from 700 MHz to 900 MHz.

Freq (MHz)

Return loss, S11 (dB)

# RF Feedthrough Measurement

Make the connection as shown in Figure 1 to determine the RF feedthrough of the mixer. RF feedthrough is a measure of the leakage of the RF power into the IF port of the mixer. Note that you must connect the synthesizer output into the LO port of the mixer via a coaxial jumper cable. Also, connect both the RF and IF ports of the mixer to the RF OUT port and RF IN port of the N9912A respectively. Keep the cable at the RF OUT port of the N9912A as short as possible.

- Launch the RF Trainer Control Panel software from the PC. Select RX Unit from the board selection and click Connect to RF Trainer. In the RX Unit Control area, click Mixer Off to power up the mixer. You will observe the LNA PCB LED light up on the training kit, and the button text change to LNA On. Next, in the Frequency Synthesizer Control area, adjust the frequency to 818 MHz. This is the LO frequency input into the mixer.

- Use the following settings:

Start		:  700 MHz

Stop		:  900 MHz

Format 	:  Log magnitude

Measurement	:  S21

Scale		:  Autoscale

N9912A setting: “[ ]”: Hardkey; “{ }”: Softkey

Set to the NA mode for S21 measurement: [Mode] > {NA} > {S21}

Set the Start Frequency to 700 MHz: [Freq/Dist] > {Start} > [700] > {MHz}

Set the Stop Frequency to 900 MHz: [Freq/Dist] > {Stop} > [900] > {MHz}

Set the output power level to LOW: [Meas Setup 4] > {Output Power} > {Low}

Set the display to the log magnitude format: [Meas Setup] > {Format} > {Log Mag}

Scale: [Scale/Amptd] > {Autoscale}

Exercises

Observations/data recording

- Find the RF feedthrough at 868 MHz.

RF feedthrough at 868 MHz = _________ dB

- Plot the RF feedthrough from the input frequency of 700 MHz to 900 MHz.

Freq (MHz)

RF Feedthrough, S21 (dB)

# LO-to-RF Isolation Measurement

LNA

Down

mixer

Frequency synthesizer

IF bandpass

filter

IF amplifier

LO In

RF Out

RF In

IF Out

Out

In

In

In

Out

Out

RF Trainer Control Panel

PC

USB cable

USB port

Receiver unit

SMA adapter

Matched load

Coaxial cable

Keep this cable as short as possible

RFA (N9912A)

RF OUT

RF IN

Figure 4 – LO-to-RF Isolation Measurement with the RFA

Make the connection as shown in Figure 4 to determine the LO-to-RF isolation. Connect both the LO In and RF In ports of the mixer to the RF OUT and RF IN ports of the N9912A respectively. The IF port of the mixer is terminated with a 50 Ω matched load. Keep the cable at the RF OUT port of the N9912A as short as possible.

- Launch the RF Trainer Control Panel software from the PC. Select RX Unit from the board selection and click Connect to RF Trainer. In the RX Unit Control area, click Mixer Off to power up the mixer. You will observe the mixer PCB LED light up on the training kit, and the button text change to Mixer On.

- Use the following settings:

Start		:  700 MHz

Stop		:  900 MHz

Format 	:  Log magnitude

Measurement	:  S21

Scale		:  Autoscale

N9912A setting: “[ ]”: Hardkey; “{ }”: Softkey

Set to the NA mode for S21 measurement: [Mode] > {NA} > {S21}

Set the Start Frequency to 700 MHz: [Freq/Dist] > {Start} > [700] > {MHz}

Set the Stop Frequency to 900 MHz: [Freq/Dist] > {Stop} > [900] > {MHz}

Set the output power level to LOW: [Meas Setup 4] > {Output Power} > {Low}

Set the display to the log magnitude format: [Meas Setup] > {Format} > {Log Mag}

Scale: [Scale/Amptd] > {Autoscale}

Exercises

Observations/data recording

- Find the LO-to-RF isolation at 818 MHz.

LO-to-RF isolation at 818 MHz = _________ dB

- Plot the isolation from the input frequency of 700 MHz to 900 MHz.

Freq (MHz)

LO−RF isolation, S21 (dB)

# LO-to-IF Isolation Measurement

Coaxial cables

Receiver unit

RFA (N9912A)

RF OUT

RF IN

RF Trainer Control Panel

PC

USB cable

USB port

Matched load

LNA

Down

mixer

Frequency synthesizer

IF bandpass

filter

IF amplifier

LO In

RF Out

RF In

IF Out

Out

In

In

In

Out

Out

SMA adapter

Figure 5 – LO-to-IF Isolation Measurement with the RFA

Make the connection as shown in Figure 5 to determine the LO-to-IF isolation. Connect both the LO In and IF Out ports of the mixer to the RF OUT and RF IN ports of the N9912A respectively. The RF port of the mixer is terminated with a 50 Ω matched load. Keep the cable at the RF OUT port of the N9912A as short as possible.

- Launch the RF Trainer Control Panel software from the PC. Select RX Unit from the board selection and click Connect to RF Trainer. In the RX Unit Control area, click Mixer Off to power up the mixer. You will observe the mixer PCB LED light up on the training kit, and the button text change to Mixer On.

- Use the following settings:

Start		:  700 MHz

Stop		:  900 MHz

Format 	:  Log magnitude

Measurement	:  S21

Scale		:  Autoscale

N9912A setting: “[ ]”: Hardkey; “{ }”: Softkey

Set to the NA mode for S21 measurement: [Mode] > {NA} > {S21}

Set the Start Frequency to 700 MHz: [Freq/Dist] > {Start} > [700] > {MHz}

Set the Stop Frequency to 900 MHz: [Freq/Dist] > {Stop} > [900] > {MHz}

Set the output power level to LOW: [Meas Setup 4] > {Output Power} > {Low}

Set the display to the log magnitude format: [Meas Setup] > {Format} > {Log Mag}

Scale: [Scale/Amptd] > {Autoscale}

Exercises

Observations/data recording

- Find the LO-to-IF isolation at 818 MHz.

LO-to-IF isolation at 818 MHz = _________ dB

- Plot the isolation from the input frequency of 700 MHz to 900 MHz.

Freq (MHz)

LO−IF isolation, S21 (dB)

References

- Application Note 1287-1, “Understanding the Fundamental Principles of Vector Network Analysis”, Keysight Technologies.

- Application Note 1287-2, “Exploring the Architectures of Network Analyzers”, Keysight Technologies.

- Product Note 8753-2, “RF Component Measurements: Mixer Measurements Using the Keysight 8753B Network Analyzer”, Keysight Technologies.

- Thomas H. Lee, “Planar Microwave Engineering”, Cambridge University Press, 2004.