ME1000 RF Circuit Design

Lab 15

Antenna Gain Measurement

with Spectrum Analyzer (SA)

____________________________________________________________________________

This courseware product contains scholarly and technical information and is protected by copyright laws and international treaties. No part of this product may be reproduced, copied, or distributed in any form or by any means without expressed written consent from Acehub Vista Sdn. Bhd.

The use of the courseware product and all other products developed and/or distributed by Acehub Vista Sdn. Bhd. are subject to the applicable License Agreement.

For further information, see Courseware Product License Agreement.

____________________________________________________________________________

Objective:

i.	To estimate the gain of an antenna via SA.

Equipments and Accessories Required:

- Spectrum analyzer

- Signal generator

- Antennas

DreamCatcherTM is the Trade Mark of Dream Catcher Consulting Sdn. Bhd.

Lab 15 – Antenna Measurement with SA

Basic Equipment Setup

CW Signal Generator

RF Output

Pt

Gt = G

Pr

Spectrum Analyzer

RF Input

Gr = G

R > Rff

Gt = Gr = G: gain of transmit and receive antenna [identical, equal gain]

Pt = Transmitted power at transmit antenna

Pr = Received power at receive antenna

Rff = far-field distance [Rff=2D2λ where D is the maximum dimension of the antenna]

Figure 1 General equipment configuration for antenna measurement using SA

Gain estimation with SA

Make the connection as shown in Figure 1 to estimate the gain of the antenna. Firstly, calculate the far-field distance, Rff. Using 868 MHz:

Wavelength @868 MHz,  = 0.3456 m (calculated using c/f )

Length of the antenna, D = 0.21 m

Supplied antenna technical specifications:

Frequency range = 806 to 960 MHz; Gain = 2.5 dBi

VSWR < 2; Polarization = Vertical; Length = 210 mm

Far-field distance, Rff = 0.2552 m (calculated using 2D2/)

- Use the following setting at the signal generator (SG):

CW Frequency 	:  868 MHz

Power		 	:  10 dBm

###### N9310A setting: “[ ]”: Hardkey; “{ }”: Softkey

Preset to default settings: [Preset]

Frequency: [Frequency] > [868] > {MHZ}

## Amplitude: [Amplitude] > [+/-] > [10] > {dBm}

Turn on output: [RF On/Off]

- Use the following setting at the SA:

Centre Frequency 	:  868 MHz

Span 		 	:  10 MHz

Input Attenuation 	:  10 dB (or auto)

RBW 			:  100 kHz (or auto)

Averaging		:  on

N9320A/B setting: “[ ]”: Hardkey; “{ }”: Softkey

Preset to default settings: [Preset/System] > {Preset}

## Center Freq: [FREQUENCY] > [868] > {MHZ}

## Span: [SPAN] > [10] > {MHZ}

## Attenuation: [AMPLITUDE] > {Attenuation} > [10] > {dB}

RBW: [BW/AVG] > {Res BW} > {100} > {KHZ}

Averaging: [BW/AVG] > {Average}

- Make sure that the distance between transmit and receive antennas, R are slightly more than the far-field distance, Rff.

OBSERVATIONS/DATA RECORDING

- What is the peak frequency and power display on the SA? Ignore the harmonic frequencies that appeared on the SA.

Peak fundamental frequency = 868 MHz

Output power level, Pr = −13 dBm

- Calculate the gain of the antenna using Friss equation as shown below:

Pr (dBm)= Gr (dB)+ Gt (dB)+ Pt (dBm) -10log4πRλ2

Gr (dB)+ Gt (dB)=Pr (dBm) - Pt (dBm)+10log4πRλ2

2G (dB)=Pr (dBm) - Pt (dBm)+10log4πRλ2, where Gr = Gt = G

= −13 dBm − 10 dBm + 25.2 dB (for R = 0.5 m)

Antenna Gain, G = 1.1 dB.

***Note: There might be some discrepancy in the observed value as compare to the expected value due to the presence of noise, reflection from ground as well as the environment where the experiment is carried out.

- Repeat the calculation at different frequencies and distance, R.

OBSERVATIONS/DATA RECORDING

- Fill in the table below.

| SG frequency, (MHz) | Wavelength  = c/f (m) | Calculated far-field distance, Rff (m) | Distance between antennas, R (m) | Measured output power at SA, Pr (dBm) | Calculated gain, G (dB) |
| --- | --- | --- | --- | --- | --- |
| 750 | 0.4 | 0.2205 | 1.5 x Rff = 0.331 | −14.8 | −2.23 |
| 750 | 0.4 | 0.2205 | 2 x Rff = 0.441 | −15.6 | −1.384 |
| 750 | 0.4 | 0.2205 | 2.5 x Rff = 0.551 | −18.2 | −1.717 |
| 868 | 0.3456 | 0.2552 | 2 x Rff = 0.51 | −13 | 1.182 |
| 868 | 0.3456 | 0.2552 | 2.5 x Rff = 0.638 | −15 | 1.155 |
| 1100 | 0.273 | 0.323 | 1.5 x Rff = 0.485 | −30.2 | −6.61 |
| 1100 | 0.273 | 0.323 | 2 x Rff = 0.646 | −32.6 | −6.567 |
| 1100 | 0.273 | 0.323 | 2.5 x Rff = 0.808 | −37.8 | −8.195 |
| 1200 | 0.25 | 0.3528 | 1.5 x Rff = 0.529 | −28.8 | −5.152 |
| 1200 | 0.25 | 0.3528 | 2 x Rff = 0.706 | −30.7 | −4.849 |
| 1200 | 0.25 | 0.3528 | 2.5 x Rff = 0.882 | −31.6 | −4.332 |
| 1300 | 0.231 | 0.382 | 1.5 x Rff = 0.573 | −30.1 | −5.112 |
| 1300 | 0.231 | 0.382 | 2 x Rff = 0.764 | −31 | −4.313 |
| 1300 | 0.231 | 0.382 | 2.5 x Rff = 0.955 | −31.2 | −3.444 |

- Is the calculated gain remains approximately the same when the distance between antennas are different? Explain your answer.

In the antenna operating frequency region, the calculated gain should be the same but there could be some discrepancies due to the reflection from the ground. Outside its operating frequency range, it will not function effectively. For more accurate measurement, anechoic chamber is essential.

Assumptions in this lab:

- Both the transmit and receive antennas are identical and have the same gain.

- Reflection is not taken into consideration when calculating the antenna gain.

References

- David M. Pozar, “Microwave Engineering”, 3rd Edition, John Wiley & Sons, 2005.