ME1000 RF Circuit Design

Lab 1

Scalar Offset Calibration with

Spectrum Analyzer (SA)

____________________________________________________________________________

This courseware product contains scholarly and technical information and is protected by copyright laws and international treaties. No part of this product may be reproduced, copied, or distributed in any form or by any means without expressed written consent from Acehub Vista Sdn. Bhd.

The use of the courseware product and all other products developed and/or distributed by Acehub Vista Sdn. Bhd. are subject to the applicable License Agreement.

For further information, see Courseware Product License Agreement.

____________________________________________________________________________

Objectives:

- To perform scalar offset calibration with spectrum analyzer.

- To verify the signal purity of a signal generator.

- To determine the losses contributed by the cables, connectors and PCB traces.

Equipments and Accessories Required:

- Spectrum analyzer

- Signal generator

- TRM calibration board

- SMA m-m coaxial cables

DreamCatcherTM is the Trade Mark of Dream Catcher Consulting Sdn. Bhd.

Lab 1 – Scalar Offset Calibration with SA

The Need for Scalar Offset Calibrations

A typical frequency domain measurement of a device is shown in Figure A1.  We usually named this device as Device-Under-Test (DUT).  When measuring the frequency response of a DUT, you will collect the power transferred from a signal generator to a spectrum analyzer via the device. To ensure that the data collected represents the actual response of the DUT only, and not portions related to the cables, connectors, and the losses within the measurement equipment itself, it is necessary to measure the total path loss of the interconnection from the equipment to the DUT’s input and output terminals. In principle, the actual effect of DUT is the measured value minus the losses due to the interconnections on both ends of the DUT.  The process of accounting for the effect of the interconnection during measurement is called “Calibration”. Since we are only interested in power measurement, the calibration process described here is known as “Scalar Offset Calibration”, and it is usually performed at various frequency of interest.

Spectrum Analyzer (SA)

RF Input

CW Signal Generator                    (SG)

RF Output

SMA coaxial cable

(Cable-SG)

Psig_gen

Pout_sa

DUT

SMA connector

SMA connector

PCB trace for connection to DUT

SMA coaxial cable

(Cable-SA)

Cable loss, Lcable_sa

Cable loss, Lcable_sg

Half-path, Lhalf_path

TRM board, Lfull_path

Half-path, Lhalf_path

The DUT can represent devices like amplifiers, filters, attenuators etc., in the trainer kits.

Figure 1 The need to calibrate scalar offset of cables, connectors and PCB trace losses

Signal generator verification and calibration of loss contributed by the cable

Spectrum Analyzer

RF Input

CW Signal Generator

RF Output

SMA coaxial cable

Psig_gen

Pout_sa

Figure 2 Signal generator verification and SMA coaxial cable loss calibration

IMPORTANT:

Excessive input power levels can damage the spectrum analyzer. The threshold for damage is different for various models. The INPUT setting can be as low as 20 dBm (0.1 Watts or 3.2 Volts* in 50 Ω load), OBSERVE the caution notice on the front panel of the equipment.

*Note: we are referring to the time-averaged power,P=V22R, where R = 50 Ω.

First, we must investigate whether the signal generator (SG) we use is a perfect linear device by connecting the SG directly to the spectrum analyzer (SA) as shown in Figure 2.

- Use the following settings for SG:

CW Frequency 	:  868 MHz

Power level		: –40dBm (will increase in 10dB step until 0dBm)

###### N9310A setting: “[ ]”: Hardkey; “{ }”: Softkey

Preset to default settings: [Preset]

Frequency: [Frequency] > [868] > {MHZ}

## Amplitude: [Amplitude] > [+/-] > [40] > {dBm}

Turn Off Mod. : [Mod On/Off]

Turn on RF Out: [RF On/Off]

- Use the following settings for the SA:

Preset the SA to its default settings.

Resolution Bandwidth	:  100 kHz

Input Attenuation 	:  10 dB (or auto)

N9320A/B setting: “[ ]”: Hardkey; “{ }”: Softkey

Preset to default settings: [Preset/System] > {Preset}

## Span: [SPAN] > {FULL SPAN}

## Attenuation: [AMPLITUDE] > {Attenuation} > [10] > {dB}

RBW: [BW/Avg] > {Res BW} > {100} > {KHZ}

OBSERVATIONS/DATA RECORDING

- What are the frequencies appear in the SA when signal power is −40dBm?

Frequency = 868 MHz

[Use marker function such as peak search if necessary]

- Increase the power level by 10dB step until 0dBm and note the different frequencies displayed. Any other frequencies appear during this power increase?

YES, they is another frequency that appears at 1740 MHz.

- Explain why there are more than single frequency component at the measured output.

The harmonic signal from the signal generator becomes high enough to be measured.

- Explain the difference between the measured and the expected output power level when the input power level is at 0 dBm.

The measured output power should be the same as the expected, but due to the increase of 2nd harmonic, some power is transfer to this harmonic, therefore, the output power will be lower.

- Fill in the table below:

| Input power from signal generator, Psig_gen (dBm) | Fundamental power (dBm) | 2nd harmonic power (dBm) | 3rd harmonic power (dBm) |
| --- | --- | --- | --- |
| −40 | −40.65 | 0 | 0 |
| −30 | −30.70 | 0 | 0 |
| −20 | −20.68 | 0 | 0 |
| −10 | −10.68 | −49.80 | 0 |
| 0 | −0.69 | −39.90 | 0 |

We can calibrate the cable losses before measuring any DUT by using the same setup as shown in Figure 2.

- Use the following settings to determine the loss contribute by the cable:

Signal Generator:

CW Frequency 	:  868 MHz

Power		 	:  −25 dBm

###### N9310A setting: “[ ]”: Hardkey; “{ }”: Softkey

Frequency: [Frequency] > [868] > {MHZ}

## Amplitude: [Amplitude] > [+/-] > [25] > {dBm}

Spectrum Analyzer:

Centre Frequency 	:  868 MHz

Span 		 	:  10 MHz (approximately 1% of centre frequency)

Input Attenuation 	:  10 dB (or auto)

RBW 			:  100 kHz (or auto)

Averaging		:  On

N9320A/B setting: “[ ]”: Hardkey; “{ }”: Softkey

## Center Freq: [FREQUENCY] > [868] > {MHZ}

## Span: [SPAN] > [10] > {MHZ}

## Attenuation: [AMPLITUDE] > {Attenuation} >[10] > {dB}

RBW: [BW/Avg] > {Res BW} > {100} > {KHZ}

Averaging: [BW/Avg] > {Average}

OBSERVATIONS/DATA RECORDING

- What is the frequency and power display on the SA when the input power level is −25dBm?

Frequency   = 868 MHz

Power level = −25.30 dBm

[Use marker function if necessary]

- Fill in the table below to determine the cable loss for various test signals. You should change the frequency and power level settings on the SG, as well as the center frequency and span settings on the SA. Use span setting as approximately 1% of centre frequency.

| SG frequency, fsig_gen (MHz) | SG power level, Psig_gen (dBm) | SA measured peak frequency, fsa (MHz) | SA measured output power, Psa (dBm) | Loss contributed by cable, Lcable = Psig_gen – Psa (dB) |
| --- | --- | --- | --- | --- |
| 50 | −25 | 50 | -25.12 | 0.12 |
| 50 | −20 | 50 | -20.2 | 0.2 |
| 50 | −15 | 50 | -15.22 | 0.22 |
| 50 | −10 | 50 | -10.26 | 0.26 |
| 50 | −5 | 50 | -5.26 | 0.26 |
| 100 | −25 | 100 | -25.32 | 0.32 |
| 100 | −10 | 100 | -10.41 | 0.41 |
| 500 | −15 | 500 | -15.77 | 0.77 |
| 500 | −10 | 500 | -15.87 | 0.87 |
| 818 | −25 | 818 | -25.92 | 0.92 |
| 818 | −20 | 818 | -20.96 | 0.96 |
| 818 | −15 | 818 | -16.02 | 1.02 |
| 818 | −10 | 818 | -11.07 | 1.07 |
| 818 | −5 | 818 | -6.058 | -1.058 |
| 818 | 0 | 818 | -1.128 | 1.128 |
| 868 | −25 | 868 | -25.92 | 0.92 |
| 868 | −20 | 868 | -20.98 | 0.98 |
| 868 | −18 | 868 | -19.08 | 1.08 |
| 868 | −15 | 868 | -16.12 | 1.12 |
| 868 | −13 | 868 | -14.12 | 1.12 |
| 868 | −10 | 868 | -11.14 | 1.14 |
| 868 | −8 | 868 | -9.132 | 1.132 |
| 868 | −5 | 868 | -6.161 | 1.161 |
| 868 | 0 | 868 | -1.218 | 1.218 |
| 1000 | −15 | 1000 | -16.18 | 1.18 |
| 1736 | −15 | 1736 | -16.58 | 1.58 |
| 1736 | −5 | 1736 | -6.618 | 1.618 |
| 2604 | −15 | 2604 | -17.04 | 2.04 |
| 2604 | −5 | 2604 | -7.022 | 2.022 |

- Does the cable loss remain the same for different input frequency?

No.

- By how much does the SA measured signal power level drop at 50 MHz and 868 MHz when the SG power level setting is reduced by 10 dB.

The SA measured signal power level drops by 0.01 to 0.04dB at 50MHz and 868MHz when the SG power level setting is reduced by 10dB.

Path Loss Calibration

##### The total path loss in a frequency response measurement consists of two parts, the input and output path loss.  Each path loss is in turn due to the coax cable and connector loss, and the loss from the copper trace on the printed circuit board (PCB) that leads to the DUT terminals.  We can measure the path loss with the use of a calibration structure called the TRM (Through-Reflect-Match) Board.

##### Input path loss: Cable-SG loss and half of the loss from the TRM board (PCB traces and connectors. Assumption: the input and output path of the TRM board are identical)

- Output path loss: Cable-SA loss and half of the loss from the TRM board.

Since both input and output path loss consists of two parts, these can be determined individually from the following procedures and summed up to give us the path loss.

Procedure:

- Measure the cable loss from Cable-SG (Lcable-sg) and Cable-SA (Lcable-sa).

- Measure the loss from the TRM board (through path).

### Equipment Setup

Spectrum Analyzer (SA)

RF Input

Signal Generator (SG)

RF Output

SMA coaxial cable

Psig_gen

Pout_sa

Required accessories

SMA m-m coaxial cable

(connecting SG)

SMA m-m coaxial cable

(connecting SA)

Figure 3 SMA coaxial cable loss measurement

We can measure the cable losses by using the setup shown in Figure 3.

- Connect the SMA coaxial cable used for connecting the SG as shown in Figure 3. This cable is referred to as Cable-SG. Label this cable for identification.

- Use the following settings for SG and SA:

SG settings:

CW Frequency 	: Frequency of interest

Power		 	: −25 dBm

###### N9310A setting: “[ ]”: Hardkey; “{ }”: Softkey

Frequency: [Frequency] > [868] > {MHZ}

## Amplitude: [Amplitude] > [+/-] > [25] > {dBm}

Turn on output: [RF On/Off]

SA settings:

Centre Frequency 	: Frequency of interest

Span 		 	: 10 MHz

Input Attenuation 	: 10 dB (or auto)

RBW 			: 100 kHz (or auto)

Averaging		: On

N9320A/B setting: “[ ]”: Hardkey; “{ }”: Softkey

## Center Freq: [FREQUENCY] > [868] > {MHZ}

## Span: [SPAN] > [10] > {MHZ}

## Attenuation: [Amplitude] > {Attenuation} > [10] > {dB}

RBW: [BW/Avg] > {Res BW} > {100} > {KHZ}

Averaging: [BW/Avg] > {Average}

Note: 50 MHz and 868 MHz are required for the other labs to compensate for the losses.

Cable loss = SA reading – SG amplitude setting

Since it is already consider as a loss, omit the negative sign of the final answer for the above equation.

- Make the measurements and record the losses for Cable-SG in Table 1:

| Frequency (MHz) | Cable-SG Loss (Lcable_sg) | Cable-SA Loss (Lcable_sa) |
| --- | --- | --- |
| 50 | 0.21 | 0.21 |
| 818 | 0.95 | 0.99 |
| 868 | 1.14 | 1.16 |
| 750 | 1.06 | 1.17 |
| 800 | 0.98 | 1.1 |
| 850 | 1.14 | 1.24 |
| 900 | 1.1 | 1.22 |
| 950 | 1.15 | 1.26 |
| 1000 | 1.18 | 1.3 |
| 1050 | 1.26 | 1.37 |
| 1100 | 1.27 | 1.39 |
| 1150 | 1.26 | 1.36 |
| 1200 | 1.32 | 1.43 |
| 1250 | 1.39 | 1.5 |
| 1300 | 1.42 | 1.52 |
| 1350 | 1.39 | 1.51 |
| 1400 | 1.43 | 1.55 |

Table 1 Path loss calibration data

- Connect the SMA coaxial cable used for connecting the SA as shown in Figure 3. This cable is referred to as Cable-SA. Label this cable for identification.

- Make the measurements and record the losses for Cable-SA in Table 1.

Losses Contributed by Connector and PCB Trace

Using a TRM board that represents the copper traces on the trainer board, we can measure the loss up till the DUT’s end.

Spectrum Analyzer (SA)

RF Input

Pout_sa

SMA connector

PCB trace for connection

to DUT

SMA coaxial cable

(Cable-SA)

Cable loss, Lcable_sa

Half-path,

Lhalf_path

TRM board  Lfull-path

TRM Board

Half-path,

Lhalf_path

Center line

CW Signal Generator        (SG)

RF Output

SMA coaxial cable

(Cable-SG)

Psig_gen

SMA connector

Figure 4 SMA connector and PCB trace loss calibration

- Make the connection as shown in Figure 4.

- The measurements from this connection are the total loss contributed by Cable-SG, Cable-SA and the full through path of the TRM board.

  - Lfull_path = Measured signal amplitude – SG amplitude setting − (Lcable_sg + Lcable_sa)

  - Lhalf_path = 0.5 * Lfull_path (Assumption: the input and output path loss are equal)

Note: The cable losses are of positive values as taken from Table 1. Omit the negative sign for Lfull_path at the end of calculation.

- Make the measurements and record the losses for the TRM Through Loss under the “Measured Power at SA” in Table 2.

- Calculate the Lfull_path and Lhalf_path and record it in Table 2.

- Calculate the total input and output path loss as below :

  - Total input path loss, Linput_path = Lcable_sg + Lhalf_path

  - Total output path loss, Loutput_path = Lcable_sa + Lhalf_path

| Frequency (MHz) | Cable-SG loss (Lcable_sg) | Cable-SA loss (Lcable_sa) | Measured Power at SA | TRM Through Loss (Lfull_path) | TRM Through Loss (Lhalf_path) | Total input path loss (Linput_path) | Total output path loss (Loutput_path) |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 50 | 0.21 | 0.21 | −25.44 | 0.02 | 0.01 | 0.22 | 0.22 |
| 818 | 0.95 | 0.99 | −27.01 | 0.07 | 0.035 | 0.985 | 1.025 |
| 868 | 1.14 | 1.16 | −27.35 | 0.05 | 0.025 | 1.165 | 1.185 |
| 750 | 1.06 | 1.17 | −27.3 | 0.07 | 0.035 | 1.095 | 1.205 |
| 800 | 0.98 | 1.1 | −27.16 | 0.08 | 0.04 | 1.02 | 1.14 |
| 850 | 1.14 | 1.24 | −27.58 | 0.2 | 0.1 | 1.24 | 1.34 |
| 900 | 1.1 | 1.22 | −27.47 | 0.15 | 0.075 | 1.175 | 1.295 |
| 950 | 1.15 | 1.26 | −27.57 | 0.16 | 0.08 | 1.23 | 1.34 |
| 1000 | 1.18 | 1.3 | −27.62 | 0.14 | 0.07 | 1.25 | 1.37 |
| 1050 | 1.26 | 1.37 | −27.88 | 0.25 | 0.125 | 1.385 | 1.495 |
| 1100 | 1.27 | 1.39 | −27.89 | 0.23 | 0.115 | 1.385 | 1.505 |
| 1150 | 1.26 | 1.36 | −27.84 | 0.22 | 0.11 | 1.37 | 1.47 |
| 1200 | 1.32 | 1.43 | −27.96 | 0.21 | 0.105 | 1.425 | 1.535 |
| 1250 | 1.39 | 1.5 | −28.14 | 0.25 | 0.125 | 1.515 | 1.625 |
| 1300 | 1.42 | 1.52 | −28.21 | 0.27 | 0.135 | 1.555 | 1.655 |
| 1350 | 1.39 | 1.51 | −28.18 | 0.28 | 0.14 | 1.53 | 1.65 |
| 1400 | 1.43 | 1.55 | −28.28 | 0.29 | 0.145 | 1.575 | 1.695 |

Table 2 Input and output path loss calibration data

References

- Presentation slides, “A Seminar on RF Measurement – Spectrum Analysis Basics”, Keysight Technologies, 2001

- Thomas H. Lee, “Planar Microwave Engineering”, Cambridge University Press, 2004.