ME1000 RF Circuit Design

Lab 8

Filter Characterization using

Spectrum Analyzer (SA)

____________________________________________________________________________

This courseware product contains scholarly and technical information and is protected by copyright laws and international treaties. No part of this product may be reproduced, copied, or distributed in any form or by any means without expressed written consent from Acehub Vista Sdn. Bhd.

The use of the courseware product and all other products developed and/or distributed by Acehub Vista Sdn. Bhd. are subject to the applicable License Agreement.

For further information, see Courseware Product License Agreement.

____________________________________________________________________________

Objective:

i.	To measure insertion loss, bandwidth, and rejection of an RF band pass filter via a SA.

Equipments and Accessories Required:

- Spectrum analyzer

- Signal generator

- ME1000 Transmitter unit

- SMA m-m coaxial cables

DreamCatcherTM is the Trade Mark of Dream Catcher Consulting Sdn. Bhd.

Note: Turn off the transceiver trainer kit when not in use. The trainer will turn off automatically when no mouse or keyboard action is detected for more than 10 minutes. Always ensure that the casing is grounded to earth and the cover is latched up before powering up the device.

Lab 8 – Filter Characterization using SA

Basic Equipment Setup

Required accessories

SMA m-m coaxial cable

Spectrum Analyzer

RF Input

CW Signal Generator

RF Output

RF band-pass filter

Power

amplifier

Up converter

Frequency synthesizer

RF Out

LO In

RF In

IF Out

In

Out

In

USB port

Transmitter unit

SMA cable

SMA Cable

Figure 1 General equipment configuration for RF filter measurement using SA

IMPORTANT:

Please use the RF Band-Pass Filter v1.01 on the RF transmitter unit for this lab measurement.

Insertion loss measurement with SA

Make the connection as shown in Figure 1 to determine the insertion loss of the RF filter. You must use the value of cable/connector/PCB trace losses, (Lcable and Lc_pcb) from previous lab.

- Use the following settings for signal generator (SG):

CW Frequency 	:  868 MHz

Power		 	:  –15 dBm

###### N9310A setting: “[ ]”: Hardkey; “{ }”: Softkey

Preset to default settings: [Preset]

Frequency: [Frequency] > [868] > {MHZ}

## Amplitude: [Amplitude] > [+/-] > [15] > {dBm}

Turn on output: [RF On/Off]

- Use the following settings for SA:

Centre Frequency 	:  868 MHz

Span 		 	:  10 MHz

Input Attenuation 	:  10 dB (or auto)

RBW 			:  100 kHz (or auto)

Averaging		:  On

N9320A/B setting: “[ ]”: Hardkey; “{ }”: Softkey

Preset to default settings: [Preset/System] > {Preset}

## Center Freq: [Frequency] > [868] > {MHZ}

## Span: [SPAN] > [10] > {MHZ}

## Attenuation: [Amplitude] > {Attenuation} > [10] > {dB}

RBW: [BW/AVG] > {Res BW} > {100} > {KHZ}

Averaging: [BW/AVG] > {Average}

OBSERVATIONS/DATA RECORDING

- What is the peak frequency and power display on the SA? Ignore the harmonic frequencies that appeared on the SA.

Peak fundamental frequency = 868 MHz

Output power level, Pout_sa_filter = −18.11 dBm

- What is the measured insertion loss of the filter when input power is –15dBm?

Gain @868MHz = Actual Filter Output power, Pout – Actual Filter Input power, Pin

= (Pout_sa_filter + Loutput-path) − (Psig_gen – Linput-path)

= (−18.11 dBm + 1.185 dB) − (−15.00 dBm − 1.165 dB)   = −0.76 dB

where,

Loutput_path = Lcable_sa + Lhalf_path (cable/connector/PCB trace losses found previously)

Linput_path = Lcable_sg + Lhalf_path (cable/connector/PCB trace losses found previously)

Insertion loss @868 MHz = − Gain @868MHz = 0.76dB

- Compare the results measured with vector network analyzer (VNA). Are they the same?

In theory, they should be the same. However, in a practical environment, the results will not be the same. VNA usually provides better accuracy due to vector correction/calibration.

- Measure the output power level of the filter at different input frequencies.

OBSERVATIONS/DATA RECORDING

- Fill in the table below. Use cable/PCB/connector losses, L (in dB) found in the previous section.

| SG frequency, fsig_gen (MHz) | Actual input power Pin = −15 − Linput_path (dBm) | Measured output power at SA, Pout_sa_filter (dBm) | Actual output power, Pout = Pout_sa_filter – Loutput_path (dBm) | Gain, G = Pout − Pin (dB) |
| --- | --- | --- | --- | --- |
| 750 | −16.03 | −30.70 | −29.57 | −13.54 |
| 800 | −16.02 | −24.76 | −23.62 | −7.60 |
| 850 | −16.04 | −19.18 | −18.04 | −2.00 |
| 900 | −16.025 | −18.32 | −17.18 | −1.15 |
| 950 | −16.07 | −18.32 | −17.14 | −1.07 |
| 1000 | −16.11 | −18.57 | −17.34 | −1.23 |
| 1050 | −16.14 | −19.64 | −18.40 | −2.26 |
| 1100 | −16.16 | −20.73 | −19.46 | −3.3 |
| 1150 | −16.18 | −20.97 | −19.70 | −3.52 |
| 1200 | −16.22 | −20.7 | −19.38 | −3.16 |
| 1250 | −16.27 | −21.55 | −20.18 | −3.91 |
| 1300 | −16.41 | −27.11 | −25.6 | −9.19 |
| 1350 | −16.39 | −32.31 | −30.8 | −14.41 |
| 1400 | −16.44 | −37.38 | −35.83 | −19.39 |

- Plot G (Gain) as a function of fsig_gen (input SG frequency) in the graph below.

- From the above plot, estimate the 3 dB bandwidth and center frequency of the RF filter.

Estimated lower 3dB frequency, fup_3db = 820 MHz

Estimated upper 3dB frequency, flow_3db = 1250 MHz

Estimated 3dB BW = fup_3db − flow_3db = 430 MHz

Centre frequency, fc = (fup_3db + flow_3db)/2 = 1035 MHz

Rejection measurement

- Use the following settings:

SG settings	:	CW Frequency 	:  500MHz

Power		 	:  –15 dBm

###### N9310A setting: “[ ]”: Hardkey; “{ }”: Softkey

Frequency: [Frequency] > [500] > {MHZ}

## Amplitude: [Amplitude] > [+/-] > [15] > {dBm}

Turn on output: [RF On/Off]

SA settings	:	Centre Frequency 	:  500MHz

Span 		 	:  10 MHz

Input Attenuation 	:  10 dB (or auto)

RBW 			:  100 kHz (or auto)

Averaging		:  On

N9320A/B setting: “[ ]”: Hardkey; “{ }”: Softkey

## Center Freq: [FREQUENCY] > [500] > {MHZ}

## Span: [SPAN] > [10] > {MHZ}

## Attenuation: [AMPLITUDE] > {Attenuation} > [10] > {dB}

RBW: [BW/AVG] > {Res BW} > {100} > {KHZ}

Averaging: [BW/AVG] > {Average}

OBSERVATIONS/DATA RECORDING

- What is the peak frequency and power display on the SA? Ignore the harmonic frequencies that appeared on the SA.

Peak fundamental frequency = 500 MHz

Output power level, Pout_sa_filter = −52.31 dBm

- What is the measured rejection at 500 MHz for the filter when input power is –15 dBm?

Rejection @500 MHz = Actual output power, Pout – Actual input power, Pin

= (Pout_sa_filter + Loutput_path) − (Psig_gen – Linput_path)

= (−52.31 dBm + 0.62 dB) − (−15.00 dBm − 0.6 dB)   = −37.05 dB

where,

Loutput_path = Lcable_sa + Lhalf_path (cable/connector/PCB trace losses found previously)

Linput_path = Lcable_sg + Lhalf_path (cable/connector/PCB trace losses found previously)

- Compare the results measured with VNA. Are they the same?

In theory, they should be the same. However, in practical environment, the results will not be the same. VNA usually provides better accuracy due to vector correction/calibration.

- Repeat for CW frequency of 1736 MHz and 2604 MHz.

Rejection @1736 MHz = (−49.56 dBm + 1.73 dB) − (−15.00 dBm − 1.69 dB) = −31.14 dB

Rejection @2604 MHz = (−44.95 dBm + 1.8 dB) − (−15.00 dBm − 1.77 dB) = −26.38 dB

References

- Presentation slides, “A Seminar on RF Measurement – Spectrum Analysis Basics”, Keysight Technologies, 2001.

- Thomas H. Lee, “Planar Microwave Engineering”, Cambridge University Press, 2004.

- David M. Pozar, “Microwave Engineering”, 3rd Edition, John Wiley & Sons, 2005.