ME1000 RF Circuit Design

Lab 1

(Solution)

Calibration with RF Analyzer

This courseware product contains scholarly and technical information and is protected by copyright laws and international treaties. No part of this publication may be reproduced by any means, be it transmitted, transcribed, photocopied, stored in a retrieval system, or translated into any language in any form, without the prior written permission of Acehub Vista Sdn. Bhd.

The use of the courseware product and all other products developed and/or distributed by Acehub Vista Sdn. Bhd. are subject to the applicable License Agreement.

For further information, see the Courseware Product License Agreement.

Objectives

- To perform scalar offset calibration with an RF analyzer (RFA)

- To determine the losses contributed by the cables, connectors, and PCB traces

Equipment Required

- Keysight N9912A Fieldfox RF Analyzer

Accessories Required

- 1 x TRM standard calibration kit

- 2 x SMA(m)-to-SMA(m) coaxial cable

- 2 x N(m)-to-SMA(f) adapter

# Introduction

The Need for Scalar Offset Calibrations

A typical frequency domain measurement of a device is shown in the figure below. This device is commonly known as Device-Under-Test (DUT). When measuring the frequency response of a DUT, you will collect the power transferred from the RFA built-in signal generator to the RFA internal receiver via the device. To ensure that the data collected represents the actual response of the DUT only, and not portions related to the cables, connectors, and the losses within the sampling circuits in the measurement equipment itself, it is necessary to measure the total path loss of the interconnection from the equipment to the DUT input and output terminals. In principle, the actual effect of DUT is the measured value minus the losses due to the interconnections on both ends of the DUT. The process of accounting for the effect of the interconnection during measurement is called “Calibration”. Since we are only interested in power measurement, the calibration process described here is known as “Scalar Offset Calibration”, and it is usually performed at various frequencies of interest.

The DUT can represent devices like amplifiers, filters, attenuators, etc., in the training kit.

Half-path (in)

Cable loss, Lcablein

Cable loss, Lcableout

SMA coaxial cable

DUT

SMA connector

SMA connector

PCB trace for connection to DUT

SMA coaxial cable

RF Analyzer

RF OUT

RF IN

Half-path (out)

Printed Circuit Board (PCB)

Figure 1 – Scalar Offset Calibration of Cables, Connectors, and PCB Trace Losses

# Scalar Calibration of Interconnection Loss

SMA coaxial cable

SMA connector

SMA connector

SMA coaxial cable

RF Analyzer

RF OUT

RF IN

The Through structure on the

TRM calibration board

Figure 2 – RF Analyzer and SMA Coaxial Cable Loss Calibration Setup

IMPORTANT:

Excessive input power levels can damage the RF analyzer. The threshold for damage is different for various models. The INPUT setting can be as low as 20 dBm (0.1 W or 3.2 V* in 50 Ω load), OBSERVE the caution notice on the front panel of the equipment. For the N9912A, the input power threshold is +27 dBm and 50 V DC.

*Note:

This refers to the time-averaged power,P=V22R, where R = 50 Ω.

Connect the RFA (N9912A) to the TRM calibration board as shown in Figure 2.

- Power up the N9912A and set it to the NA (Network Analyzer) mode, then select the S21 measurement. The N9912A display is as shown in . Here assuming we are interested in the frequency range from 2 MHz to 3 GHz.

###### N9912A setting: “[ ]” – Hardkey (The keypad on the instrument)

“{}” – Softkey (The keys underneath the display)

Preset to default settings: [Preset] > {Preset}

Set to the NA mode: [Mode] > {NA} > {S21}

Set the frequency range: [Freq/Dist] > {Stop} > [3] > {GHZ}

Set the output power level to Low: [Meas Setup 4] > {Output Power} > {LOW}

Note:

Setting the output power level at the RF OUT port to high will saturate active circuits like the amplifier or mixer. Using a low power setting will be sufficient for most measurements.

Losses due to cables, connectors, and internal sampling circuitry of the N9912A

0 dB

Figure 3 – N9912A Uncalibrated S21 Magnitude vs. Frequency

- There are TWO approaches to account for the path loss: (1) by manually using the waveform math functions to perform simple scalar calibrations, or (2) by using the Normalize function in the Calibration Menu. The first method will be demonstrated first. Begin by storing the uncalibrated data in the RFA internal memory, then employ the waveform math function to account for losses of the cables and connectors. The calibrated S21 magnitude measurement is shown in . This approach is useful for equipment that does not support the built-in calibration function.

###### N9912A setting: “[ ]”: Hardkey, “{}”: Softkey

Store the initial data in memory: [Trace 6] > {Data > Mem}

Normalize the waveform: {Data Math} > {Data / Mem}

Normalized |S21|

Marker

Marker readings

Figure 4 – N9912A Calibrated S21 Magnitude vs. Frequency (Manual Calibration Using Waveform Math)

Note:

You can remove the waveform math function by going through the following sequence:

[Trace 6] > {Data Math} > {Math Off}

- We can also invoke the Normalize function in the Calibration Menu. The mechanism is similar to those outlined in step 2, except the procedures are automated by the instrument. The calibrated S21 magnitude is shown in . Note the Cal ON icon in the screenshot.

###### N9912A setting: “[ ]”: Hardkey, “{}”: Softkey

Engage calibration: [Cal 5] > {Calibration Type} > {Normalize}

Start calibration: {Start Cal} > {Thru} > {Measure} > {Finish}

Note the Cal ON icon

Figure 5 – N9912A Calibrated S21 Magnitude vs. Frequency (Using the Normalize Function)

Note:

Use the marker ([Marker] hardkey) to record  your readings. The marker is used to obtain a reading at a particular frequency or amplitude. For simplicity, you may wish to use the Marker Function or Marker Search option to search for maximum and minimum points.

Reference

- “Keysight Fieldfox RF Analyzer N9912A User’s Guide”, Keysight Technologies, 2009.