ME1000 RF Circuit Design

Lab 4

(Solution)

Filter Characterization Using RF Analyzer

This courseware product contains scholarly and technical information and is protected by copyright laws and international treaties. No part of this publication may be reproduced by any means, be it transmitted, transcribed, photocopied, stored in a retrieval system, or translated into any language in any form, without the prior written permission of Acehub Vista Sdn. Bhd.

The use of the courseware product and all other products developed and/or distributed by Acehub Vista Sdn. Bhd. are subject to the applicable License Agreement.

For further information, see the Courseware Product License Agreement.

Objectives

- To characterize the RF filter with reflection and transmission measurement via an RF analyzer (RFA)

- To display results in multiple plots such as the Smith Chart and Magnitude Plot

Equipment Required

- ME1000 Transmitter unit

- Keysight N9912A FieldFox RF Analyzer

Accessories Required

- 2 x SMA(m)-to-SMA(m) coaxial cable

IMPORTANT:

Turn off the training kit when not in use. The kit will turn off automatically when no mouse or keyboard action is detected for more than 10 minutes. Always ensure that the casing is grounded and the cover is latched up before powering up the device.

# Introduction

Basic Equipment Setup

RF Analyzer

RF OUT

RF IN

RF bandpass filter

Power

amplifier

Upconverter

Frequency synthesizer

RF Out

LO In

RF In

IF Out

In

Out

Out

In

USB port

Transmitter unit

SMA cable

SMA cable

Figure 1 – General Equipment Configuration for Filter Measurement

The RF OUT port of the RF analyzer (RFA) generates a sinusoidal wave. The wave interacts with the device-under-test (DUT) and the result is captured by the RF IN port of the RFA. Therefore, the RF OUT port is always connected to the input of the DUT and the RF IN port is connected to the output of the DUT.

In this lab, you need to set the RFA to the Network Analyzer (NA) mode. When performing a transmission, or an S21 measurement, carry out the calibration procedures from Lab 1 with the SMA coaxial cables connected to both the RF OUT port and RF IN port of the RFA before you proceed. The cable and connectors on the RF OUT port of the RFA should be as short as possible and of high quality.

Note:

If your RFA comes with Option 111 (‘QuickCal’), you will be able to perform external calibration on the RF OUT port for reflection or an S11 measurement. Refer to the Keysight N9912A User’s Guide for the detailed procedure.

When viewing signals, the [Freq/Dist], [Scale/Amptd], and [Marker] hardkeys are the most helpful. When the [Scale/Amptd] hardkey is pressed, the {Autoscale} option can be quite handy, but you will rarely use it as you get more comfortable with manipulating the display.

Note: 
Use the marker ([Marker] hardkey) to record your readings. The marker is used to obtain a reading at a particular frequency or amplitude. For simplicity, you may wish to use the Marker Function or Marker Search option to search for maximum and minimum points.

# Transmission Measurements

- Set the RFA (N9912A) to the NA mode for S21 measurement.

- Perform the Normalize calibration as in Lab 1 from 2 MHz to 4 GHz with the LOW output power setting.

- Use the following settings to determine the insertion loss of the filter:

Format 	:  Log magnitude

Measurement	:  S21

Scale		:  Autoscale

N9912A setting: “[ ]”: Hardkey; “{ }”: Softkey

Set to the NA mode for S21 measurement: [Mode] > {NA} > {S21}

Set the Start Frequency to 2 MHz: [Freq/Dist] > {Start} > [2] > {MHz}

Set the Stop Frequency to 4 GHz: [Freq/Dist] > {Stop} > [4] > {GHz}

Set the output power level to LOW: [Meas Setup 4] > {Output Power} > {Low}

Scale: [Scale/Amptd] > {Autoscale}

Perform Calibration: Normalization

- Connect the RF bandpass filter as shown in Figure 1.

Exercises

Observations/data recording – General

- What is the filter type? (bandpass, low pass, high pass, or not sure)

Band-pass filter.

Use the marker function to find the following parameters:

3dB cutoff frequencies (use minimum loss as the reference value):

Lower 3dB frequency (fL) = 825 MHz

Upper 3dB frequency (fH) = 1250 MHz

3dB bandwidth = 425 MHz

Centre frequency (fc) = 1037.5 MHz

[Note: Centre frequency is defined as mean of the lower and upper 3dB frequencies]

Insertion loss (S21) at centre frequency = −2.5 dB

[Note: The insertion loss should be −3 dB or less at center frequency]

- What is the importance of insertion loss?

A small insertion loss in the passband minimizes unnecessary signal attenuation. Moreover, the insertion contributes to increase in noise figure if this filter is used in a wireless receiver.

- What is the importance of insertion loss flatness?

The insertion loss flatness indicates the linearity of the filter. A very linear filter should have a better flatness.

- Explain why the S21 magnitude increases slowly in the upper stopband with frequency.

This is called RF leakage, could be due to (1) parasitic capacitance, e.g. electric field coupling between the Input and Output transmission lines via the surrounding metallic structures. (2) the properties of the discrete components which deteriorates with frequency.

Observations/data recording – Out-of-band rejection

- Find the rejection at stopband (at 300 MHz, 500 MHz, and 800 MHz offset from the center frequency).

Upper rejection @300 MHz = −14.1 dB

Lower rejection @300 MHz = −15 dB

Upper rejection @500 MHz = −29.6 dB

Lower rejection @500 MHz = −37 dB

Upper rejection @800 MHz = −38 dB

Lower rejection @800 MHz = −57.8 dB

# Reflection Measurements

- Set the N9912A to the NA mode for S11 measurement.

- Use the following settings to determine the insertion loss of the filter:

Format 	:  Log magnitude

Measurement	:  S11

Scale		:  Autoscale

N9912A setting: “[ ]”: Hardkey; “{ }”: Softkey

Set to the NA mode for S21 measurement: [Mode] > {NA} > {S11}

Set the display to the Log magnitude format: [Meas Setup] > {Format} > {LogMag}

Scale: [Scale/Amptd] > {Auto Scale}

Exercises

Observations/data recording

- What is the return loss at pass-band?

Return loss, S11 at pass-band = 13 dB at 868 MHz with the highest return loss of 23.5 dB occurs at 930 MHz.

- What is the return loss at stop-band?

Return loss, S11 at stop-band = 1 dB (at frequencies lower than 700 MHz and at frequencies higher than 1.4 GHz)

[Note: You can choose any frequency point within the pass-band or stop-band]

- What is the importance of return loss at the pass-band?

Return loss indicates the quality of the matching of the filter at pass-band with respect to the source. A large return loss (> 10 dB) is required for good matching at pass-band. A small return loss is required at stop-band for good rejection.

- From the return loss observed at pass-band and stop-band, describe the filter’s matching at both bands.

Pass-band – Good matching condition (all frequencies could pass)

Stop-band – Total mismatch condition (all frequencies rejected)

- Change the settings to display the Smith Chart for impedance measurement.

- From the Smith Chart, you can determine the magnitude and phase of the reflection coefficient. From these values, you can convert the reflection coefficient into the complex form, and finally convert the reflection coefficient into the input impedance Zin, 
via Zin=1+Γ111−Γ11⋅Zo , where Zo is 50 

Format 	:  Smith Chart

N9912A setting: “[ ]”: Hardkey; “{ }”: Softkey

Set the display to the Smith Chart format: [Meas Setup] > {Format} > {Smith}

Exercises

Observations/data recording

- What is the impedance at passband?

Impedance at passband = 54.4 – j 24.4 Ω at 868 MHz, 56.8 + j 1.35 Ω at 930 MHz.

- What is the impedance at stopband?

Impedance at stopband = 5.49 – j 0.232 Ω

The impedance at stopband varies as it follows the edge of the Smith Chart. At different points in the Smith Chart plot, the impedance varies from having high values to low values.

[Note: You can choose any frequency point within the passband or stopband]

- What is the ideal impedance at passband?

Ideal Impedance at passband = 50 Ω

- What is the ideal impedance at stopband?

Ideal impedance at the stopband should be 0 Ω or anything that differs from c) by a huge amount.

References

- Application Note 1287-1, “Understanding the Fundamental Principles of Vector Network Analysis”, Keysight Technologies.

- Application Note 1287-2, “Exploring the Architectures of Network Analyzers”, Keysight Technologies.

- Application Note 1287-4, “Network Analyzer Measurements: Filter and Amplifier Examples”, Keysight Technologies.

- “Keysight FieldFox RF Analyzer N9912A User’s Guide”, Keysight Technologies, 2009.

- Thomas H. Lee, “Planar Microwave Engineering”, Cambridge University Press, 2004.