## Slide 1
The University of British Columbia, Vancouver, British Columbia, Canada
PredictUS: A Method to Extend the Resolution Precision Trade-off in Quantitative Ultrasound Image Reconstruction
Farah Deeba, Robert Rohling

## Slide 2: Attenuation Coefficient Estimate (ACE)
A promising clinical tool to detect and monitor fatty liver;
C:\k-WAVE\k-Wave\miccai\IUS\IUS Oral
Fat Fraction (%)
ACE
Deeba, F., et al.: SWTV-ACE: spatially weighted regularization based attenuation
coefficient estimation method for hepatic steatosis detection. In: International Conference on Medical Image Computing and Computer-Assisted Intervention (2019)

## Slide 3: Attenuation Coefficient Estimate (ACE)
C:\k-WAVE\k-Wave\miccai\IUS\IUS Oral
A promising clinical tool to detect and monitor fatty liver
Potential for placental tissue characterization;
Fetal
Maternal
ACE
Deeba, F., et al.: Attenuation coefficient estimation of normal placentas. Ultrasound Med. Biol. 45(5), 1081–1093 (2019)

## Slide 4: ACE Computation
= Power Spectral Density (PSD) Term
𝒙=[𝛼 𝛽];
= Attenuation Coefficient Estimate
= Backscatter Coefficient
= System Matrix

## Slide 5: ACE Computation: The Fundamental Trade-off
PSD (dB)
Frequency (MHz)
Estimated PSD
2 cm
PSD Estimation
Ultrasound RF Data

## Slide 6: ACE Computation: The Fundamental Trade-off
C:\k-WAVE\k-Wave\miccai\IUS\IUS Oral
ACE Error
Window Size = 2 mm
Window Size = 3 mm
Axial Window Size (mm)
Window Size = 4 mm
Window Size = 5 mm
Window Size = 6 mm
Window Size = 7 mm
Window Size = 8 mm
Spectral based Method
2 cm
Resolution
Accuracy & Precision

## Slide 7: Why not choose the largest window?
Poor resolution will limit the clinical uptake for
Heterogeneous tissue (e.g. placenta);
Thin tissue (e.g. dermis with 1-2 mm thickness)
C:\k-WAVE\k-Wave\miccai\IUS\IUS Oral

## Slide 8: Proposed Solution: PredictUS
Axial Window Size =2 mm
(100 RF samples)
100 RF samples (original)
50 RF samples (predicted)

## Slide 9: WaveNet Model based Sequence-to-Sequence Convolution Neural Network
Linear
Training Stage
1-step ahead prediction
Testing Stage
n-step ahead prediction

## Slide 10: Receptive Field
Dilation Factors,

## Slide 11: Data and Experiment
Training examples: 106,496 data segments (ACE = 0.1 dB/cm/MHz - 1.65 dB/cm/MHz).
Test examples: 550 data segments (ACE = 0.1 dB/cm/MHz - 1.5 dB/cm/MHz)
Notes: Our proposed RNN setup is able to accurately learn how to model normal traffic. Figure 3 illustrates the average loss value of the training and validation dataset during the process of learning. It is worth noticing that the training loss is higher than the validation loss, contrary to intuition, as the loss function for validation does not use regularization and the learning process uses three times more data, resulting in higher average loss values. After each batch of training, the weights of the network are adjusted and the average loss of all batches is reported. The RNN is able to learn, in only few epochs, an accurate representation of normality, as can be appreciated from the low loss in the validation data.

## Slide 12: Result: Ultrasound RF Data Prediction
Input Predicted Target

## Slide 13: Result: Ultrasound RF Data Prediction
n =16
n =15 step prediction

## Slide 14: Result: ACE Computation
For resolution comparison, we define an axial dimension of 10λ to be the optimum.
Liu, W., Zagzebski, J.A.: Trade-offs in data acquisition and processing parameters for backscatter and scatterer size estimations. IEEE Trans. Ultrason. Ferroelectr. Freq. Control 57(2), 340–352 (2010)
We consider following three cases with same axial size (10λ) with varying input and predicted data length. The input data length for the three cases are:
  Case I: 7.5λ
  Case II: 5λ
  Case III: 2.5λ

## Slide 15: Result: ACE Computation
Small Window PredictUS Window Large Window
Input Predicted Target
Case I: Input window Length 7.5λ

## Slide 16: Result: ACE Computation
Small Window PredictUS Window Large Window
Input Predicted Target
Case II: Input window Length 5λ

## Slide 17: Result: ACE Computation
Small Window PredictUS Window Large Window
Input Predicted Target
Case III: Input window Length 2.5λ

## Slide 18: Conclusion
We conduct a simulation based proof-of-study to extend the trade-off between resolution and ACE measurement quality.
A resolution improvement up to 50% (input axial dimension was reduced from 5λ from the optimal requirement of 10λ) was attained with comparable measurement quality.
Future work includes extension of the method to other Quantitative Ultrasound parameters and application on physical phantom and biological tissue.

## Slide 19: Appendix

## Slide 20: Training: 1-step Ahead Prediction
Encode
Predict
Encode
Predict
Training
Validation
Walk-Forward Split Mode

## Slide 21: Testing: n-step Ahead Prediction
Encode
Predict
Testing
Encode
Predict
Encode
Predict
…